|  | The 2nd International Conference (Computational Mechanics and Virtual Engineering( COMEC 2007 11 – 13 OCTOBER 2007, Brasov, Romania |
| --- | --- |

on Mechanical Valve Prosthesis

and Its Quality Evaluation Parameters

Ileana ROSCA 1

1Transilvania University, Brasov, ROMANIA,  ilcrosca@unitbv.ro

Abstract: Cardiovascular illness treatment demand extremely various solutions from simple drugs till supporting circulation temporary devices for implanted artificial hearts. The problem of valve prosthesis quality parameters is a very sensitive one, because it must ensure not only functional but also physiological. In this paper we propose to present the dependence between valve’s diameter and its type, patient’s homodynamic characteristics and its physical activity. Also, in order to choose the appropriate diameter of a valve prosthesis one must take into account the relationship between pressure passing through valve, the temporary evolution of the blood volume and the cardiac frequency. Further, some dimensional and non-dimensional parameters that estimate the valve’s quality function are presented: the effective passing section, the effective free section, and reflux.

Keywords: heart, valve, blood pressure, diameter, quality evaluation

1. INTRODUCTION

Many ill human parts need mechanical or electric devices: kidney, blood and vessels, valves, heart's muscle etc. Cholesterol particles and grease deposed internally to thicken blood vessels drive flow reduction, reduced oxygenation till occlusion. If the plates are deposed on valves (Figure 1), they can produce heart attack. Growing of cholesterol plates is foreseeing by angiographies, and other tests.

Figure 1: Human heart

Generally the drugs are not sufficient and their mechanical removal or vessel’s enlargement by graft or by-pass is imposed (even relapses are often reported).

Valve’s principal illnesses (Figure 2)are insufficiency (regurgitant - incapacity to stop retrogressive flow during closing phase) and incapacity to open enough (stenotic - due to partial welding of valve’s boards). Both illnesses can coexist and diagnosis is quite rapid and simple through specific symptoms, utilized intervention being plastic valves’ surgery or, in grave cases, replacement of valves with mechanical or biological devices.

Figure 2: Valvular functions.

2. Prothtetic heart problems

Valve prosthesis are of two types: biologic and mechanic. Mechanic valves are done from materials having a higher strength than the natural ones tissues (figure 3) as special steels and are doing themselves the guidance of blood flow. But regularity in flow running is not always ensured (and also coagulation), with great deviations, turbulent and slow zones, presenting great danger for blood as thrombosis.

This last mentioned phenomenon is aggravated by closing mechanism of valve leafs which have the tendency to crush red blood cells. In any case, mechanical valve prosthesis a much long life lasting, effort resistant and can be used for all types of replacement (aortic, mitral, tricuspid).

(a) 			(b)			          (c)			       (d)

Figure 3: Mechanic valves: Starr-Edwards (a); single leafelet valve Medtronic-Hall (b); bileafelet valve St. Jude (c) Hancock valve (d).

Biologic valves are composed from three parts: biologic tissue, tissue for fixation and support ring (stent) and made by cutting leafs from animal tissues: pig valve, bovine pericardial, and human tissue, mounted on metallic, polymeric or biologic support.

Preparing biologic valves supposes many phases and preliminary treatment to reduce rejection process and, consequently, they are very expansive. This valve type rarely present rapid and fatal damage – it appears gradually and provoke specific symptoms.

(a)				         (b)				  (c)

Figure 4: Biologic valves: basic structure (a); porcine valve (b); human tissue valve (c)

3. CALCULUS OF VALVE PROSTHESIS OPTIMAL DIAMETER

Calculus of optimal diameter of valve prosthesis to be implanted is done in accord with valve type (aortic or mitral), patient hemodynamic characteristics and physical activity usually performed (repose or effort). For a correct choice of valve, it is necessary to take into account the connections between pressure drop at its end, temporally evolution of the volume and cardiac frequency.

Valves can be considered as zones with dimensional variation producing instantaneous energy loss, illustrated by pressure variation, given by the next equation:

with ,	    (1)

where: ( is energetic loss coefficient, q is blood flow, (  the blood density, D the passing diameter in valve and k is defined as a constant proportional with energy loss at the end of the prosthesis.

Average energetic loss during the period T (ventricular systole or diastole) is given by the integral of instantaneous energetic losses and this drives us to the conclusion that it is necessary to know the temporary evolution of volume. For the aortic valve blood volume (cardiac discharges) the Swanson and Clark (1977) equation is representative (eq. 2) and for volume evolution through mitral valve, the Talukder and Reul (1978) equation (3).

, 	 (2)

,		(3)

where: Ts indicates the duration of ejection phase (ventricular systole), Td is the duration of ventricular diastole, Qs is the maximal value of instantaneous aortic volume and Qm the same value for mitral volume (recharge or ventricular diastole).

If the duration of the whole cardiac cycle is defined as: Tc=Td+Ts=f -1 the average volume will be given by:

(4)

After transformations in equation (2), (3) and (4), the instantaneous systolic and diastolic maxim volume is:

,		(5)

with a relationship between Td and Ts calculated by Katz and Feil (1923): , the constant 0,096 being measured in seconds. To augment cardiac frequency, the systolic phase is quite constant and diastole is reduced.

Using previous equations, average pressure variation (energetic loss) in aortic valve is determined as function of frequency f average volume, passing blood diameter D and the constant k of energetic loss. Analogous, the same calculus will be performed for mitral valve, obtaining the next two equations:

,	(6)

with the values: 12,88 expressed in seconds-1 and 1,51, non dimensional. Logically, the above equations (6) allow to obtain the inner diameters of valves:

(7)

Using previous relations, abacus allowing the rapid choose of the inner diameter of the prosthesis to be implanted are done, in accord with the patient characteristics (average volume, frequency, transvalvular pressure variation) and valve type (mitral or aortic), indicated by k parameter.

4. QUALITY evaluation parameters FOR A mechanic valvular prosthesis

In order to evaluate the quality of functional behavior, several parameters are more often used, as:

- EOA (effective orifice area), used by cardiologist as to know the efficient passing section:

[cm2],		(8)

with Q [l/min] – blood volume and ΔPPF [mm(Hg] – maximal transvalvular pressure variation.

- DC (discharging coefficient), is a non dimensional coefficient expressing the ratio between EOA (an equivalent section) and inner area of the prosthesis (a measured area, representing the effective free area for blood pass):

(9)

- PI (performance index) is still a non dimensional parameter given by the ration between EOA and mounting surface Am of the prosthesis, meaning that it shows the effect of joint (welding):

.	(10)

If DC is kept constant, a greater PI indicates a greater reduction of the diameter due to the valve’s joint and, reciprocally, if PI is constant, a greater DC allows a better blood flow (bigger section) and les (P.

- RF (ebbing percentage) represents the percentage of volume passing through the closed valve (regurgitant working):

. 	(11)

5. Conclusions

In vascular pathology, valvular illnesses still present a great incidence. In their treatment, especially for advanced valvular damages, it is strongly necessary to replace the natural one with an artificial valve. From the hydraulic point of view, a heart valve is a mono directional one that guides the blood flow by opening and closing itself and by producing pressure gradients. Apparently, very simple in fact it has a complex structure and a correlated function with other functions of intra cardiac structures. Mechanisms of valves’ damaging by different pathologic process are still not completely known and also, details that confer to it the mechanical durability as alive structure.

REFERENCES

- Bignardi C., Sesana Raffaella. Corso di biomecanica, Torino, 1997,

- Denischi A., Marin I., Antonescu D., Petrescu R. Biomecanica, Ed. Academiei, Bucuresti, 1989,

- Groza, P. Fiziologie umana, Editura Medicala, Bucuresti. 1980,

- Griffith M., Lamancusa J. S. Rapid Prototyping Technologies, Rapid Prototyping. 1998,

- Olariu V., Rosca I. and others. Biomechanics, Ed. Macarie, Targoviste, Romania, 1988.